Question: An alcohol-based surgical rub requires that hands are

(a) dried with a sterile towel before gloving

(b) wet from the tap to activate the rub

(c) clean and dry before application, rubbed over all surfaces per protocol and allowed to dry completely before gloving

(d) rinsed with water after rub before gloving

Answer: c

Solution: ✅ Correct Answer: c. Clean and dry before application, rubbed over all surfaces per protocol and allowed to dry completely before gloving.

🧠 Clinical Rationale: Rubs need dry skin and full coverage, and hands must dry before gloves else wetness traps flora ⏳.

🔍 Key Clues: Dry skin, full rub, air dry.

❌ Why Not Others: (b) water dilutes the product; (a) towel drying removes the active layer; (d) rinsing after rub removes the protective residue.

⚡ 5-Second Recall: Rub dry, glove dry 🎯

📚 Ref: WHO surgical hand preparation.

Positive Marks: 1

Negative Marks: 0.33

Question: A sterility indicator on a pack shows a failed colour change. The correct action is

(a) add a second indicator strip to the pack

(b) treat the pack as non-sterile, remove it from circulation and reprocess or discard it

(c) use the pack since the autoclave was serviced recently

(d) store it at the back of the shelf for emergencies

Answer: b

Solution: ✅ Correct Answer: b. Treat the pack as non-sterile, remove it from circulation and reprocess or discard it.

🧠 Clinical Rationale: Failed indicators mean the process was not demonstrated, so the pack cannot be released for patient use 🚫.

🔍 Key Clues: Failed indicator: no release.

❌ Why Not Others: (c) recent servicing does not validate this load; (a) adding indicators after the cycle proves nothing; (d) storage does not restore sterility.

⚡ 5-Second Recall: No indicator, no release 🎯

📚 Ref: CSSD release criteria.

Positive Marks: 1

Negative Marks: 0.33

Question: The central sterile supply department (CSSD) workflow follows the sequence

(a) receipt, decontamination, inspection and packing, sterilisation, sterile storage and issue

(b) storage, issue, receipt, decontamination, sterilisation

(c) sterilisation, receipt, packing, issue, decontamination

(d) packing, receipt, sterilisation, decontamination, storage

Answer: a

Solution: ✅ Correct Answer: a. Receipt, decontamination, inspection and packing, sterilisation, sterile storage and issue.

🧠 Clinical Rationale: The workflow moves one way from dirty to clean, and each stage protects the next in the chain 🏭.

🔍 Key Clues: Dirty first, sterile last.

❌ Why Not Others: (d), (c) and (b) place cleaning after or mixed with sterile handling, reversing the safe flow.

⚡ 5-Second Recall: Receipt to issue, one way 🎯

📚 Ref: CSSD workflow standards; NABH.

Positive Marks: 1

Negative Marks: 0.33

Question: Needles should never be

(a) counted if a needle is missing

(b) disposed of in a sharps container immediately

(c) discarded with the syringe intact into the sharps bin

(d) recapped by hand, bent, broken or removed from syringes before disposal

Answer: d

Solution: ✅ Correct Answer: d. Recapped by hand, bent, broken or removed from syringes before disposal.

🧠 Clinical Rationale: Recapping, bending and removal are the classic mechanisms of needlestick injury for theatre and waste staff 🚫.

🔍 Key Clues: No recap, no bend, no remove.

❌ Why Not Others: (b) immediate sharps disposal is correct practice; (a) tracking missing needles is required rather than prohibited; (c) intact syringe disposal into the sharps bin is accepted.

⚡ 5-Second Recall: Handle freely, then bin it 🎯

📚 Ref: sharps injury prevention guidance.

Positive Marks: 1

Negative Marks: 0.33

Question: Masks in theatre should be changed between cases or when

(a) a new surgeon arrives

(b) they become damp, soiled or contaminated, since moisture compromises filtration

(c) the nurse feels like a fresh mask

(d) the room is cleaned

Answer: b

Solution: ✅ Correct Answer: b. They become damp, soiled or contaminated, since moisture compromises filtration.

🧠 Clinical Rationale: Moisture bridges the mask fibres and lets organisms pass, so wet masks are replaced 😷.

🔍 Key Clues: Wet mask: swap it.

❌ Why Not Others: (a) staffing changes do not dictate mask change; (c) preference is not a criterion; (d) room cleaning does not renew a mask.

⚡ 5-Second Recall: Damp mask: replace 🎯

📚 Ref: AORN attire guideline.

Positive Marks: 1

Negative Marks: 0.33

Question: Which item should NOT be sterilised in a steam autoclave?

(a) cotton drapes

(b) stainless steel instruments

(c) glassware

(d) oils and greases in sealed containers

Answer: d

Solution: ✅ Correct Answer: d. Oils and greases in sealed containers.

🧠 Clinical Rationale: Steam cannot penetrate oil or powder or reach sealed containers, so these items require dry heat or other methods 🛢️.

🔍 Key Clues: Oil and powder do not take steam.

❌ Why Not Others: (b), (a) and (c) are all standard steam-load items.

⚡ 5-Second Recall: No steam through oil 🎯

📚 Ref: sterilisation method selection guidance.

Positive Marks: 1

Negative Marks: 0.33

Question: Loan instrument sets arriving from a vendor require

(a) storage in the dirty area until needed

(b) documented receipt, inspection for cleanliness and function, and reprocessing according to policy before use

(c) direct use since the vendor states sterility

(d) reuse of vendor packaging without checks

Answer: b

Solution: ✅ Correct Answer: b. Documented receipt, inspection for cleanliness and function, and reprocessing according to policy before use.

🧠 Clinical Rationale: Vendor-processed sets have caused outbreaks, so hospital-controlled reprocessing and inspection are mandatory 🔍.

🔍 Key Clues: Loan sets get our process.

❌ Why Not Others: (c) vendor claims are not accepted without hospital verification; (d) packaging is inspected, and sets are reprocessed; (a) dirty storage contaminates sets.

⚡ 5-Second Recall: Loan sets: inspect and reprocess 🎯

📚 Ref: loan instrument guidance; infection control.

Positive Marks: 1

Negative Marks: 0.33

Question: Radioactive waste from theatre must be handled by

(a) the hospital laundry

(b) the biomedical waste contractor in a yellow bag

(c) security staff in a general bin

(d) the approved radioactive waste agency under regulatory rules, stored in shielded marked containers and never in routine waste streams

Answer: d

Solution: ✅ Correct Answer: d. The approved radioactive waste agency under regulatory rules, stored in shielded marked containers and never in routine waste streams.

🧠 Clinical Rationale: Radioactive waste follows atomic energy regulations, kept separate from biomedical waste under shielded, marked storage ☢️.

🔍 Key Clues: Radioactive: a separate regulator.

❌ Why Not Others: (b), (c) and (a) all breach radiation waste rules and endanger handlers.

⚡ 5-Second Recall: Radioactive waste: separate stream 🎯

📚 Ref: AERB rules; BMW Rules 2016.

Positive Marks: 1

Negative Marks: 0.33

Question: The most important safety step after applying an alcohol-based skin prep is to

(a) drape immediately while wet for better adhesion

(b) blot it dry with sterile gauze

(c) allow it to dry completely before draping or using diathermy to prevent fire

(d) add more alcohol for extra effect

Answer: c

Solution: ✅ Correct Answer: c. Allow it to dry completely before draping or using diathermy to prevent fire.

🧠 Clinical Rationale: Alcohol vapour trapped under drapes ignites with diathermy, causing surgical fires, so full drying is mandatory 🔥.

🔍 Key Clues: Dry before drape and cautery.

❌ Why Not Others: (a) wet draping traps flammable vapour; (b) blotting removes the active prep; (d) more alcohol increases fire risk.

⚡ 5-Second Recall: Let the alcohol dry 🎯

📚 Ref: surgical fire safety guidance; AORN.

Positive Marks: 1

Negative Marks: 0.33

Question: Physical monitoring of the autoclave involves checking

(a) the mood of the operator

(b) the number of packs in the load

(c) time, temperature, pressure and cycle charts or printouts for each load

(d) only the colour of the indicator tape

Answer: c

Solution: ✅ Correct Answer: c. Time, temperature, pressure and cycle charts or printouts for each load.

🧠 Clinical Rationale: Physical parameters are the first-line check that the cycle ran within its specified envelope 📈.

🔍 Key Clues: Time, temperature, pressure logged.

❌ Why Not Others: (d) tape is a chemical indicator; (b) load size does not prove the cycle; (a) operator mood is not data.

⚡ 5-Second Recall: Physical monitor: time, temp, pressure 🎯

📚 Ref: CSSD monitoring standards.

Positive Marks: 1

Negative Marks: 0.33

Question: Battery-powered surgical instruments are managed by

(a) discarding after one use in all cases

(b) sterilising with batteries fully charged and installed always

(c) washing in water with batteries in place

(d) removing or checking batteries per manufacturer instructions, cleaning contacts and confirming function before sterilisation

Answer: d

Solution: ✅ Correct Answer: d. Removing or checking batteries per manufacturer instructions, cleaning contacts and confirming function before sterilisation.

🧠 Clinical Rationale: Battery compartments and contacts have specific cleaning and sterilisation rules, and function is verified before release 🔋.

🔍 Key Clues: Follow the device instructions.

❌ Why Not Others: (b) leaving batteries in may void validity; (c) water immersion damages electronics; (a) most powered instruments are reusable with care.

⚡ 5-Second Recall: Check the battery rules 🎯

📚 Ref: manufacturer instructions; CSSD standards.

Positive Marks: 1

Negative Marks: 0.33

Question: Placenta or anatomical waste handed over to a family, where permitted on request, requires

(a) no documentation as it is a personal matter

(b) following the statutory conditions and infection control precautions, with the request documented

(c) disposal in the general waste bin after the handover

(d) use of the yellow bag for transfer to the family

Answer: b

Solution: ✅ Correct Answer: b. Following the statutory conditions and infection control precautions, with the request documented.

🧠 Clinical Rationale: Handing anatomical material to families is permitted under defined conditions, and documentation protects both parties 📋.

🔍 Key Clues: Statutory conditions plus documentation.

❌ Why Not Others: (a) undocumented handovers breach the rules; (c) general waste disposal of anatomical material is prohibited; (d) yellow bags are not used for family handover.

⚡ 5-Second Recall: Anatomical handover: document it 🎯

📚 Ref: BMW Rules 2016 provisions.

Positive Marks: 1

Negative Marks: 0.33

Question: The single best summary of surgical asepsis is

(a) gloves alone suffice for sterile procedures

(b) only the surgeon needs aseptic technique

(c) hand preparation, correct gowning and gloving, and disciplined sterile field management with immediate action on any break

(d) sterility lasts once achieved during a case

Answer: c

Solution: ✅ Correct Answer: c. Hand preparation, correct gowning and gloving, and disciplined sterile field management with immediate action on any break.

🧠 Clinical Rationale: Surgical asepsis is a continuous chain: preparation, donning, field discipline and rapid correction of breaches ✔️.

🔍 Key Clues: Prep, don, guard, correct.

❌ Why Not Others: (a) gloves without technique fail; (d) sterility breaks and must be repaired; (b) the whole team maintains asepsis.

⚡ 5-Second Recall: Asepsis is a chain, not a moment 🎯

📚 Ref: AORN sterile technique; WHO SSI prevention.

Positive Marks: 1

Negative Marks: 0.33

Question: Autoclave chamber cleaning and maintenance are important because

(a) maintenance has no impact on sterilisation

(b) they change the aesthetic appearance

(c) cleaning is only needed annually

(d) scale, debris and worn gaskets cause wet loads, air leaks and cycle failures

Answer: d

Solution: ✅ Correct Answer: d. Scale, debris and worn gaskets cause wet loads, air leaks and cycle failures.

🧠 Clinical Rationale: Chamber condition, door seals and drains determine whether steam reaches its parameters reliably 🛠️.

🔍 Key Clues: Clean chamber, tight seal, dry loads.

❌ Why Not Others: (b) appearance is irrelevant to function; (c) scheduled maintenance is more frequent than annual; (a) maintenance directly affects sterilisation outcomes.

⚡ 5-Second Recall: Gaskets and drains matter 🎯

📚 Ref: steriliser maintenance standards.

Positive Marks: 1

Negative Marks: 0.33

Question: Instrument tracking systems using barcodes or RFID mainly enhance

(a) traceability of which instruments were in which set, on which cycle and, where relevant, to which patient

(b) the colour coding of trays

(c) the strength of the instruments

(d) the surgeon's operating speed

Answer: a

Solution: ✅ Correct Answer: a. Traceability of which instruments were in which set, on which cycle and, where relevant, to which patient.

🧠 Clinical Rationale: Digital tracking makes recalls fast and accurate, replacing paper guesswork 📡.

🔍 Key Clues: Track every tray, every cycle.

❌ Why Not Others: (c) tracking does not change instrument strength; (d) speed is a clinical factor; (b) colour coding is one optional aid, not the system purpose.

⚡ 5-Second Recall: Tracking makes recall possible 🎯

📚 Ref: instrument traceability standards.

Positive Marks: 1

Negative Marks: 0.33

Question: The most common cause of biomedical waste rules violations in hospitals is

(a) mixing of waste categories at the point of generation because of poor segregation practice

(b) over-training of staff

(c) too many colour-coded bins

(d) use of closed containers

Answer: a

Solution: ✅ Correct Answer: a. Mixing of waste categories at the point of generation because of poor segregation practice.

🧠 Clinical Rationale: Segregation is the discipline that the whole system rests on, and mixing is the leading violation 🗑️.

🔍 Key Clues: Mixing is the top failure.

❌ Why Not Others: (c) adequate colour-coded bins help rather than harm; (b) training is protective; (d) closed containers are required practice.

⚡ 5-Second Recall: Most violations start with mixing 🎯

📚 Ref: BMW compliance reviews.

Positive Marks: 1

Negative Marks: 0.33

Question: If sterile saline soaks through a sterile drape onto the table beneath, the drape is

(a) still sterile because saline is sterile

(b) allowed if covered with another towel

(c) considered contaminated by strike-through and must be replaced

(d) acceptable until the case ends

Answer: c

Solution: ✅ Correct Answer: c. Considered contaminated by strike-through and must be replaced.

🧠 Clinical Rationale: Fluid carries organisms from the unsterile table through the drape, so wet-through drapes are replaced 🚱.

🔍 Key Clues: Wet means strike-through, replace.

❌ Why Not Others: (a) the underlying surface is unsterile regardless of the fluid; (b) covering does not undo contamination; (d) contamination lasts the whole case.

⚡ 5-Second Recall: Strike-through equals replace 🎯

📚 Ref: AORN sterile technique guideline.

Positive Marks: 1

Negative Marks: 0.33

Question: According to the Spaulding classification, critical items are those that

(a) touch intact mucous membranes only

(b) contact intact skin only

(c) touch the floor of the theatre

(d) enter sterile tissue or the vascular system and must be sterilised

Answer: d

Solution: ✅ Correct Answer: d. Enter sterile tissue or the vascular system and must be sterilised.

🧠 Clinical Rationale: Critical items such as instruments, implants and catheters carry the highest infection consequence, so they demand sterilisation 🔪.

🔍 Key Clues: Critical: enters tissue, sterilise.

❌ Why Not Others: (a) mucous membrane contact defines semi-critical, needing high-level disinfection; (b) skin contact defines non-critical; (c) floor contact is environmental, not Spaulding critical.

⚡ 5-Second Recall: Critical: sterilise 🎯

📚 Ref: CDC Spaulding classification.

Positive Marks: 1

Negative Marks: 0.33

Question: Sterilisation records should be retained

(a) until the paper is full

(b) never, since records are optional

(c) for one week only

(d) for the period required by national policy so that traceability can be reconstructed years later

Answer: d

Solution: ✅ Correct Answer: d. For the period required by national policy so that traceability can be reconstructed years later.

🧠 Clinical Rationale: Implant-related infections can present late, so records are kept for the mandated retention period 🗄️.

🔍 Key Clues: Keep records for years.

❌ Why Not Others: (c) a week is far too short; (a) until the paper is full is not a policy; (b) records are mandatory.

⚡ 5-Second Recall: Records outlive the case 🎯

📚 Ref: CSSD documentation standards.

Positive Marks: 1

Negative Marks: 0.33

Question: When managing a spill, the person cleaning should first

(a) mop immediately with the nearest mop

(b) leave it for housekeeping to discover

(c) place a warning sign, don PPE, and contain the spill with the absorbent material from the spill kit

(d) wipe with bare hands for speed

Answer: c

Solution: ✅ Correct Answer: c. Place a warning sign, don PPE, and contain the spill with the absorbent material from the spill kit.

🧠 Clinical Rationale: Signage plus PPE protect others while containing the spill limits spread before disinfection ⚠️.

🔍 Key Clues: Sign, PPE, contain.

❌ Why Not Others: (a) immediate mopping spreads the spill and risks falls; (d) bare hands risk exposure; (b) unattended spills expose staff and patients.

⚡ 5-Second Recall: Warn, protect, contain 🎯

📚 Ref: spill management protocol.

Positive Marks: 1

Negative Marks: 0.33

Question: Double gloving is recommended for

(a) high-risk procedures with sharp instruments or known blood-borne infection risk

(b) situations where gloves are unavailable

(c) every minor procedure

(d) cases where the surgeon requests it socially

Answer: a

Solution: ✅ Correct Answer: a. High-risk procedures with sharp instruments or known blood-borne infection risk.

🧠 Clinical Rationale: Double gloving reduces perforation and blood contact risk in orthopaedic, trauma and other high-risk surgery 🧤.

🔍 Key Clues: High risk: two gloves.

❌ Why Not Others: (c) routine minor cases do not require it; (d) social preference is not a clinical indication; (b) unavailability is not an indication.

⚡ 5-Second Recall: Sharps risk: double glove 🎯

📚 Ref: AORN sharps safety; CDC guidance.

Positive Marks: 1

Negative Marks: 0.33

Question: The hot air oven achieves sterilisation by

(a) ultraviolet light exposure

(b) dry heat at approximately 160 degrees Celsius for 2 hours using conduction and oxidation

(c) ethylene oxide gas exposure

(d) steam under pressure at 134 degrees

Answer: b

Solution: ✅ Correct Answer: b. Dry heat at approximately 160 degrees Celsius for 2 hours using conduction and oxidation.

🧠 Clinical Rationale: Dry heat kills by oxidation and protein denaturation, suiting glass, metal and powders that steam cannot penetrate 🔥.

🔍 Key Clues: Dry heat: 160 for 2 hours.

❌ Why Not Others: (d) steam is a different method entirely; (c) ethylene oxide is a chemical gas method; (a) ultraviolet light is a surface adjunct, not a steriliser.

⚡ 5-Second Recall: Hot air: 160, 2 hours 🎯

📚 Ref: CDC sterilisation guidance.

Positive Marks: 1

Negative Marks: 0.33

Question: CSSD staff should receive hepatitis B vaccination and health surveillance because

(a) occupational exposure to blood and body fluids is inherent to instrument reprocessing

(b) exposure does not occur in CSSD

(c) vaccination is only for theatre staff

(d) surveillance replaces PPE

Answer: a

Solution: ✅ Correct Answer: a. Occupational exposure to blood and body fluids is inherent to instrument reprocessing.

🧠 Clinical Rationale: Dirty-zone work carries sharps and splash exposure, so vaccination and health monitoring protect staff 💉.

🔍 Key Clues: Protect the reprocessing team.

❌ Why Not Others: (c) all exposed staff are covered; (d) surveillance complements, never replaces, PPE; (b) exposure is a documented CSSD hazard.

⚡ 5-Second Recall: Vaccinate the dirty-zone team 🎯

📚 Ref: occupational health standards.

Positive Marks: 1

Negative Marks: 0.33

Question: Soiled dressings, cotton and linen contaminated with blood from an infected patient should go into

(a) white sharps containers

(b) red bags with tubing

(c) yellow bags because they are soiled with blood or body fluids

(d) the general waste bag used for the ward

Answer: c

Solution: ✅ Correct Answer: c. Yellow bags because they are soiled with blood or body fluids.

🧠 Clinical Rationale: Blood-soaked dressings and linen belong to the solid soiled category, which is the yellow stream 🩸.

🔍 Key Clues: Blood-soaked items: yellow.

❌ Why Not Others: (b) red is for contaminated plastics and tubing; (a) white is for sharps; (d) general waste cannot receive infectious dressings.

⚡ 5-Second Recall: Soiled means yellow 🎯

📚 Ref: BMW Management Rules 2016.

Positive Marks: 1

Negative Marks: 0.33

Question: During the surgical hand scrub, hands must be kept

(a) lower than the elbows to drain the water away

(b) level with the elbows while scrubbing

(c) position is not important during scrub

(d) higher than the elbows throughout so water flows from clean hands towards the elbows

Answer: d

Solution: ✅ Correct Answer: d. Higher than the elbows throughout so water flows from clean hands towards the elbows.

🧠 Clinical Rationale: Gravity keeps rinse water moving from the cleanest area (hands) to the least clean (elbows) 🚰.

🔍 Key Clues: Fingers up, elbows down.

❌ Why Not Others: (a) hands below elbows lets dirty water flow back over the fingers; (b) level positioning loses the gradient; (c) hand position is a scrub principle.

⚡ 5-Second Recall: Hands above elbows 🎯

📚 Ref: WHO surgical hand preparation.

Positive Marks: 1

Negative Marks: 0.33

Question: A limitation of glutaraldehyde use in practice is that it

(a) is only effective on metal instruments

(b) cannot kill bacteria at all

(c) irritates skin, eyes and airways and requires ventilation and PPE during handling

(d) requires no monitoring of concentration

Answer: c

Solution: ✅ Correct Answer: c. Irritates skin, eyes and airways and requires ventilation and PPE during handling.

🧠 Clinical Rationale: Glutaraldehyde is a potent irritant and sensitiser, so controlled ventilation, PPE and monitoring of solution strength are essential 🧯.

🔍 Key Clues: Strong irritant: ventilate and protect.

❌ Why Not Others: (b) it is broadly antimicrobial including mycobacteria; (d) concentration must be monitored; (a) it works on many material types.

⚡ 5-Second Recall: Glutaraldehyde: protect the airway 🎯

📚 Ref: CDC high-level disinfection; occupational safety guidance.

Positive Marks: 1

Negative Marks: 0.33

Question: Hand hygiene facilities in CSSD should be

(a) replaced by gloves alone

(b) handwash sinks at zone boundaries and alcohol rubs available at workstations, with staff trained in the five moments

(c) present but not used because gloves protect

(d) available only in the staff toilet

Answer: b

Solution: ✅ Correct Answer: b. Handwash sinks at zone boundaries and alcohol rubs available at workstations, with staff trained in the five moments.

🧠 Clinical Rationale: Zone transitions and task changes are hygiene moments, and gloves never replace hand hygiene 🧼.

🔍 Key Clues: Sinks at boundaries, rub at benches.

❌ Why Not Others: (d) distant sinks are not used at transitions; (a) gloves leak and are removed, so hygiene remains; (c) gloves do not eliminate the need for hand hygiene.

⚡ 5-Second Recall: Hygiene at every zone change 🎯

📚 Ref: WHO hand hygiene; CSSD standards.

Positive Marks: 1

Negative Marks: 0.33

Question: Displaying the correct colour-coded bins in theatre without training would most likely result in

(a) continued segregation errors, showing that training and monitoring accompany infrastructure

(b) automatic rule compliance

(c) perfect compliance

(d) no change in practice

Answer: a

Solution: ✅ Correct Answer: a. Continued segregation errors, showing that training and monitoring accompany infrastructure.

🧠 Clinical Rationale: Infrastructure enables compliance, but only training, supervision and feedback change practice 📚.

🔍 Key Clues: Bins alone are not enough.

❌ Why Not Others: (c) perfect compliance needs trained staff; (d) practice can worsen without instruction; (b) compliance is behavioural, not automatic.

⚡ 5-Second Recall: Train, then the bins work 🎯

📚 Ref: BMW implementation experience; training guidance.

Positive Marks: 1

Negative Marks: 0.33

Question: Medical asepsis differs from surgical asepsis because medical asepsis

(a) requires sterile gowns for all contacts

(b) sterilises all equipment before use

(c) reduces the number and spread of microorganisms using clean technique rather than eliminating them

(d) is used only in the operating theatre

Answer: c

Solution: ✅ Correct Answer: c. Reduces the number and spread of microorganisms using clean technique rather than eliminating them.

🧠 Clinical Rationale: Medical asepsis (clean technique) limits transmission, whereas surgical asepsis demands sterility 🧼.

🔍 Key Clues: Clean reduces, sterile eliminates.

❌ Why Not Others: (b) sterilisation is a surgical asepsis concept; (a) sterile gowns belong to surgical settings; (d) medical asepsis is used throughout the hospital.

⚡ 5-Second Recall: Clean reduces, sterile removes 🎯

📚 Ref: infection control texts; AORN guidance.

Positive Marks: 1

Negative Marks: 0.33

Question: The single best summary of sterilisation practice is

(a) clean thoroughly, choose the method that suits the device, monitor every load and act on failures with recall

(b) reuse single-use items when stock is low

(c) rely on chemical indicators alone

(d) autoclave everything at maximum temperature

Answer: a

Solution: ✅ Correct Answer: a. Clean thoroughly, choose the method that suits the device, monitor every load and act on failures with recall.

🧠 Clinical Rationale: Sterilisation quality is a chain: cleaning, method choice, monitoring and failure response, each protecting the next 🔗.

🔍 Key Clues: Clean, match, monitor, recall.

❌ Why Not Others: (d) not all devices tolerate maximum heat; (c) indicators do not prove sterility; (b) reuse of single-use devices is unsafe and non-compliant.

⚡ 5-Second Recall: Clean, match, monitor, recall 🎯

📚 Ref: CDC sterilisation; CSSD standards.

Positive Marks: 1

Negative Marks: 0.33

Question: A pack found to be damp on the sterile store shelf should be

(a) left in place with a note for the next shift

(b) dried in the sun and returned

(c) removed from stock, investigated for storage or cycle causes and reprocessed

(d) used first to prevent waste

Answer: c

Solution: ✅ Correct Answer: c. Removed from stock, investigated for storage or cycle causes and reprocessed.

🧠 Clinical Rationale: Dampness indicates barrier failure from storage or cycle faults, so the pack is removed and the cause sought 🌫️.

🔍 Key Clues: Damp pack: remove, investigate.

❌ Why Not Others: (b) drying cannot restore sterility; (d) using wet packs risks infection; (a) passing it on prolongs the hazard.

⚡ 5-Second Recall: Wet pack leaves the shelf 🎯

📚 Ref: CSSD storage standards.

Positive Marks: 1

Negative Marks: 0.33

Question: The contact time of a disinfectant is important because

(a) organisms are killed only when the wet surface remains exposed for the required contact duration

(b) wiping dry immediately improves efficacy

(c) contact time applies only to fogging

(d) longer contact always bleaches surfaces

Answer: a

Solution: ✅ Correct Answer: a. Organisms are killed only when the wet surface remains exposed for the required contact duration.

🧠 Clinical Rationale: Disinfectant labels specify wet contact times; wiping dry early shortens killing and wastes the agent ⏳.

🔍 Key Clues: Keep it wet for the time stated.

❌ Why Not Others: (d) bleaching is not the concern; (c) the principle applies to all chemical disinfection; (b) drying early reduces effect.

⚡ 5-Second Recall: Respect the contact time 🎯

📚 Ref: CDC disinfection guidance.

Positive Marks: 1

Negative Marks: 0.33

Question: Sterile items are added to a sterile field only when

(a) the package is intact, dry and in date, and the item is transferred without contacting unsterile surfaces

(b) the wrapper is damp but sealed

(c) the item is needed urgently regardless of packaging

(d) the package looks old but unopened

Answer: a

Solution: ✅ Correct Answer: a. The package is intact, dry and in date, and the item is transferred without contacting unsterile surfaces.

🧠 Clinical Rationale: Package integrity, dryness and expiry are the sterility indicators, and the transfer method protects the item 📦.

🔍 Key Clues: Intact, dry, in date.

❌ Why Not Others: (d) age matters through expiry dates; (c) urgency does not sterilise damaged packs; (b) moisture breaches the barrier.

⚡ 5-Second Recall: Check pack before adding 🎯

📚 Ref: AORN sterile technique; CSSD standards.

Positive Marks: 1

Negative Marks: 0.33

Question: Sterilisation is best defined as

(a) the reduction of organisms to a safe level

(b) the removal of visible dirt from instruments

(c) the killing of vegetative bacteria only

(d) the complete destruction or removal of all microorganisms including bacterial spores

Answer: d

Solution: ✅ Correct Answer: d. The complete destruction or removal of all microorganisms including bacterial spores.

🧠 Clinical Rationale: Sterilisation is absolute and includes spores; disinfection reduces organisms but does not guarantee spore removal 🧪.

🔍 Key Clues: Sterile: spores gone too.

❌ Why Not Others: (a) it is a reduction concept (disinfection/antisepsis); (b) cleaning removes soil but not all organisms; (c) spore killing distinguishes sterilisation.

⚡ 5-Second Recall: Sterile means spores gone 🎯

📚 Ref: CDC disinfection and sterilisation guideline; Spaulding classification.

Positive Marks: 1

Negative Marks: 0.33

Question: The correct recall procedure after a sterilisation failure begins with

(a) identifying affected loads through traceability records, notifying users, retrieving items and reprocessing, then documenting the investigation

(b) discarding only the failed test strip

(c) informing only the CSSD manager

(d) waiting for a patient infection to appear

Answer: a

Solution: ✅ Correct Answer: a. Identifying affected loads through traceability records, notifying users, retrieving items and reprocessing, then documenting the investigation.

🧠 Clinical Rationale: Recall speed depends on traceability, with immediate retrieval and clear communication to all users 🚨.

🔍 Key Clues: Find, tell, retrieve, document.

❌ Why Not Others: (d) waiting for infections is a failure of governance; (b) the test strip is not the problem; (c) all affected users must know.

⚡ 5-Second Recall: Recall: trace, notify, retrieve 🎯

📚 Ref: CSSD recall standards; NABH.

Positive Marks: 1

Negative Marks: 0.33

Question: Incineration is the treatment of choice for

(a) general hospital waste

(b) red category recyclable plastics

(c) yellow category waste including anatomical and soiled items, at a facility meeting emission standards

(d) sharps containers only

Answer: c

Solution: ✅ Correct Answer: c. Yellow category waste including anatomical and soiled items, at a facility meeting emission standards.

🧠 Clinical Rationale: High-temperature incineration destroys anatomical and soiled waste completely, with emission controls protecting the environment 🔥.

🔍 Key Clues: Yellow burns.

❌ Why Not Others: (b) red plastics are autoclaved and recycled; (d) sharps are sterilised and shredded; (a) general waste is not incinerated as biomedical waste.

⚡ 5-Second Recall: Incinerate the yellow stream 🎯

📚 Ref: BMW Rules 2016 treatment standards.

Positive Marks: 1

Negative Marks: 0.33

Question: Sterile technique during minor procedures such as central line insertion requires

(a) a sterile field only if the patient requests it

(b) full surgical asepsis with cap, mask, sterile gown and gloves, full-body drape and skin antisepsis

(c) alcohol wipe of the gloves before insertion

(d) clean gloves and a small drape only

Answer: b

Solution: ✅ Correct Answer: b. Full surgical asepsis with cap, mask, sterile gown and gloves, full-body drape and skin antisepsis.

🧠 Clinical Rationale: Inserting a device into a deep sterile site demands maximal sterile barrier precautions to prevent bloodstream infection 🩺.

🔍 Key Clues: Lines need full barriers.

❌ Why Not Others: (d) partial barriers raise infection risk; (c) glove wiping is not asepsis; (a) patient preference does not govern technique.

⚡ 5-Second Recall: Line insertion: maximal barriers 🎯

📚 Ref: CDC CLABSI prevention; AORN guidance.

Positive Marks: 1

Negative Marks: 0.33

Question: Wet loads in the autoclave most often result from

(a) correct loading with proper spacing

(b) overloading, poor drying time, wet steam or blocked drains

(c) cool instruments entering the chamber

(d) excessive drying time

Answer: b

Solution: ✅ Correct Answer: b. Overloading, poor drying time, wet steam or blocked drains.

🧠 Clinical Rationale: Wet packs signal a process fault such as poor steam quality, drainage problems or inadequate drying 💧.

🔍 Key Clues: Wet load: find the fault.

❌ Why Not Others: (a) correct loading prevents wet loads; (c) cool items do not cause wet loads; (d) long drying prevents moisture.

⚡ 5-Second Recall: Wet packs mean a fault 🎯

📚 Ref: CSSD troubleshooting guidance.

Positive Marks: 1

Negative Marks: 0.33

Question: Autoclaves in CSSD should be positioned

(a) in the sterile store's open shelving area

(b) between the clean packing area and the sterile store so loads move forward without crossing dirty zones

(c) inside the decontamination area

(d) in the corridor for ventilation

Answer: b

Solution: ✅ Correct Answer: b. Between the clean packing area and the sterile store so loads move forward without crossing dirty zones.

🧠 Clinical Rationale: Placement supports the unidirectional flow, loading from the clean side and unloading toward sterile storage 🏭.

🔍 Key Clues: Load clean, unload sterile.

❌ Why Not Others: (c) dirty-zone sterilising risks recontamination; (d) corridor placement blocks traffic and flow; (a) sterilisers do not sit inside storage racks.

⚡ 5-Second Recall: Steriliser between clean and sterile 🎯

📚 Ref: CSSD design standards.

Positive Marks: 1

Negative Marks: 0.33

Question: The maximum storage time for biomedical waste before treatment, as a general rule, is

(a) no limit if the bags are sealed

(b) 7 days at room temperature

(c) 48 hours

(d) 1 month

Answer: c

Solution: ✅ Correct Answer: c. 48 hours.

🧠 Clinical Rationale: Waste held beyond 48 hours degrades, smells and multiplies organisms, so defined storage limits apply ⏰.

🔍 Key Clues: Two days maximum.

❌ Why Not Others: (b) a week exceeds the standard limit; (d) a month is far beyond any permitted period; (a) sealed bags still deteriorate and grow organisms.

⚡ 5-Second Recall: Store waste under 48 hours 🎯

📚 Ref: BMW Management Rules 2016.

Positive Marks: 1

Negative Marks: 0.33

Question: When a scrubbed person must step away from the table briefly, they should

(a) step back facing the field, keep hands high and return without touching non-sterile surfaces

(b) turn their back to the field and walk away

(c) remove one glove to move freely

(d) lower their hands below the waist

Answer: a

Solution: ✅ Correct Answer: a. Step back facing the field, keep hands high and return without touching non-sterile surfaces.

🧠 Clinical Rationale: Facing the field and keeping hands high preserves both the person's sterility and the field's integrity 🚶.

🔍 Key Clues: Back away facing forward.

❌ Why Not Others: (b) turning the back risks contaminated contact; (d) hands below waist are considered unsterile; (c) removing gloves ends scrubbed status.

⚡ 5-Second Recall: Back away, keep facing 🎯

📚 Ref: AORN sterile technique guideline.

Positive Marks: 1

Negative Marks: 0.33

Question: Staff loading an autoclave should verify that

(a) the load is as full as possible

(b) towels are stuffed into containers for weight

(c) indicators are expired for realism

(d) packs are dry, correctly wrapped, labelled with the cycle, and the load matches the cycle type selected

Answer: d

Solution: ✅ Correct Answer: d. Packs are dry, correctly wrapped, labelled with the cycle, and the load matches the cycle type selected.

🧠 Clinical Rationale: Loading discipline plus cycle matching prevents failed loads, and every pack carries traceability labelling 🔍.

🔍 Key Clues: Right cycle, right load, labelled.

❌ Why Not Others: (a) overloading causes failures; (c) expired indicators are invalid; (b) packing materials into containers blocks steam.

⚡ 5-Second Recall: Check the load before the cycle 🎯

📚 Ref: CSSD practice standards.

Positive Marks: 1

Negative Marks: 0.33

Question: The three-sink method in manual decontamination is used to

(a) soak instruments for three hours

(b) dry instruments in three stages

(c) clean three sets at once for efficiency

(d) wash, rinse and then disinfect or lubricate instruments in separate sequential sinks to avoid recontamination

Answer: d

Solution: ✅ Correct Answer: d. Wash, rinse and then disinfect or lubricate instruments in separate sequential sinks to avoid recontamination.

🧠 Clinical Rationale: Separate stages stop dirty water from contacting already-cleaned items, completing cleaning before disinfection 🚰.

🔍 Key Clues: Wash, rinse, disinfect: three sinks.

❌ Why Not Others: (a) soaking time is not the method; (c) simultaneous processing recontaminates; (b) drying is a separate unheated step.

⚡ 5-Second Recall: Three sinks, three stages 🎯

📚 Ref: CSSD manual cleaning guidance.

Positive Marks: 1

Negative Marks: 0.33

Question: The most effective sequence for cleaning soiled theatre surfaces is

(a) in a random order chosen by the cleaner

(b) clean to dirty, working from cleaner areas outward and top to bottom to avoid recontamination

(c) only the visibly soiled spots

(d) dirty to clean, finishing at the cleanest area

Answer: b

Solution: ✅ Correct Answer: b. Clean to dirty, working from cleaner areas outward and top to bottom to avoid recontamination.

🧠 Clinical Rationale: Direction and sequence protect cleaned surfaces from becoming re-soiled during the same task 🧭.

🔍 Key Clues: Clean to dirty, top to bottom.

❌ Why Not Others: (d) finishing on clean surfaces with a dirty cloth recontaminates them; (a) random order risks recontamination; (c) visible-only cleaning leaves organisms behind.

⚡ 5-Second Recall: Clean outward, clean downward 🎯

📚 Ref: environmental cleaning standards.

Positive Marks: 1

Negative Marks: 0.33

Question: Before beginning the surgical scrub, the nurse should

(a) apply hand cream for smoothness

(b) remove jewellery, watches and nail polish, keep nails short and clean, and cover hair and mask

(c) keep wristwatch on to time the scrub

(d) wear rings for grip on instruments

Answer: b

Solution: ✅ Correct Answer: b. Remove jewellery, watches and nail polish, keep nails short and clean, and cover hair and mask.

🧠 Clinical Rationale: Jewellery, polish and long nails harbour organisms and prevent thorough cleaning of hands 💍.

🔍 Key Clues: Bare hands, short nails before scrub.

❌ Why Not Others: (a) cream after scrub is optional policy but never before; (d) rings are prohibited; (c) wall clocks or timers replace wristwatches.

⚡ 5-Second Recall: Strip hands before scrub 🎯

📚 Ref: WHO hand hygiene; AORN attire standards.

Positive Marks: 1

Negative Marks: 0.33

Question: A laryngoscope blade contacts mucous membranes during use and is therefore classified as

(a) a non-critical item needing only wiping

(b) a semi-critical item requiring at least high-level disinfection, with full sterilisation preferred

(c) a critical item because it touches the airway

(d) an item requiring no processing

Answer: b

Solution: ✅ Correct Answer: b. A semi-critical item requiring at least high-level disinfection, with full sterilisation preferred.

🧠 Clinical Rationale: Mucous membrane contact puts the blade in the semi-critical class, and many units sterilise blades for added safety 🩺.

🔍 Key Clues: Mucous membrane: semi-critical.

❌ Why Not Others: (a) intact-skin items are non-critical; (c) critical applies to sterile tissue or vascular contact; (d) all patient-contact items are processed.

⚡ 5-Second Recall: Blade: semi-critical 🎯

📚 Ref: Spaulding classification; CDC.

Positive Marks: 1

Negative Marks: 0.33

Question: Air pressure in the CSSD packing area is typically maintained

(a) positive relative to the decontamination area so air moves from clean to dirty

(b) negative, drawing dirty air towards packing

(c) neutral with open doors for balance

(d) equal to the dirty area pressure

Answer: a

Solution: ✅ Correct Answer: a. Positive relative to the decontamination area so air moves from clean to dirty.

🧠 Clinical Rationale: Directional airflow protects packed items by keeping contaminated air on the dirty side 🌬️.

🔍 Key Clues: Clean pushes air to dirty.

❌ Why Not Others: (b) negative packing pressure pulls contamination toward clean stock; (d) equal pressure allows drift both ways; (c) open doors break any gradient.

⚡ 5-Second Recall: Packing side under positive pressure 🎯

📚 Ref: CSSD air handling standards.

Positive Marks: 1

Negative Marks: 0.33

Question: Between-case cleaning of the operating room is best described as

(a) full terminal cleaning of every surface including the ceiling

(b) spraying air freshener for odour control

(c) cleaning only the floor and nothing else

(d) cleaning surfaces that came into contact with the previous patient and the immediate environment, using detergent and disinfectant per policy

Answer: d

Solution: ✅ Correct Answer: d. Cleaning surfaces that came into contact with the previous patient and the immediate environment, using detergent and disinfectant per policy.

🧠 Clinical Rationale: Between-case cleaning breaks transmission from case to case within the list, focusing on contact surfaces 🧽.

🔍 Key Clues: Case-to-case: contact surfaces.

❌ Why Not Others: (c) floors alone miss the table, rails and equipment; (a) that is terminal cleaning; (b) deodorising does not disinfect.

⚡ 5-Second Recall: Clean contact surfaces each case 🎯

📚 Ref: CDC environmental cleaning guidance.

Positive Marks: 1

Negative Marks: 0.33

Question: A scrubbed nurse must never

(a) turn their back on or reach across the sterile field

(b) ask the circulator for supplies

(c) keep arms above waist level

(d) face the sterile field while working

Answer: a

Solution: ✅ Correct Answer: a. Turn their back on or reach across the sterile field.

🧠 Clinical Rationale: The sterile person's back and clothing edges are unsterile, and reaching across passes over sterile items dangerously 🚫.

🔍 Key Clues: Face the field, never cross it.

❌ Why Not Others: (d) facing the field is required; (c) high arms maintain sterility; (b) requesting supplies from the circulator is normal practice.

⚡ 5-Second Recall: No back turned, no reaching over 🎯

📚 Ref: AORN sterile technique guideline.

Positive Marks: 1

Negative Marks: 0.33

Question: Sterile services staff competency matters because

(a) errors are invisible and harmless

(b) competency checks are optional paperwork

(c) only doctors need training

(d) processing errors such as wrong loading, wrong cycles or missed indicators directly threaten every patient in the hospital

Answer: d

Solution: ✅ Correct Answer: d. Processing errors such as wrong loading, wrong cycles or missed indicators directly threaten every patient in the hospital.

🧠 Clinical Rationale: CSSD errors multiply across all users, so documented training and competency checks are a hospital-wide safety control 🎓.

🔍 Key Clues: One error, many patients.

❌ Why Not Others: (c) technicians and nurses perform the processing; (b) competency is a required control; (a) errors cause infections and failed recalls.

⚡ 5-Second Recall: CSSD errors hit everyone 🎯

📚 Ref: CSSD competency standards.

Positive Marks: 1

Negative Marks: 0.33