Question: Before sterilisation, hinged and ratcheted instruments should be

(a) locked closed to protect the jaws

(b) wrapped individually with the ratchet closed

(c) opened so the box lock, hinge and serrations are exposed to cleaning and steam

(d) stacked tightly in the tray

Answer: c

Solution: ✅ Correct Answer: c. Opened so the box lock, hinge and serrations are exposed to cleaning and steam.

🧠 Clinical Rationale: Closed hinges harbour soil and block steam, so instruments are opened and sometimes held with a pin 🔓.

🔍 Key Clues: Open the hinges to clean and steam.

❌ Why Not Others: (a) closed jaws hide contamination; (d) tight stacking prevents steam and drying; (b) individual wrapping with closed ratchets still traps soil.

⚡ 5-Second Recall: Open hinges, open to steam 🎯

📚 Ref: CSSD instrument preparation.

Positive Marks: 1

Negative Marks: 0.33

Question: Allis forceps are typically used to

(a) clamp a bleeding vessel deep in the pelvis

(b) hold a needle during suturing

(c) grasp tissue such as fascia, skin and rectus sheath using their interlocking teeth

(d) retract the rib cage

Answer: c

Solution: ✅ Correct Answer: c. Grasp tissue such as fascia, skin and rectus sheath using their interlocking teeth.

🧠 Clinical Rationale: Allis jaws have interlocking teeth for secure gripping of tough layers requiring traction 👌.

🔍 Key Clues: Allis grips fascia and skin.

❌ Why Not Others: (a) deep vessels need long haemostats; (b) needle holders hold needles; (d) rib retraction uses specialised retractors.

⚡ 5-Second Recall: Allis for fascia 🎯

📚 Ref: instrument atlases.

Positive Marks: 1

Negative Marks: 0.33

Question: When laying out an instrument table, heavier and larger instruments are placed

(a) only on the left side of the table regardless of handedness

(b) on top of the delicate instruments to save space

(c) in a separate room to reduce clutter

(d) at the back or lower level away from delicate instruments, with frequently used items within easy reach

Answer: d

Solution: ✅ Correct Answer: d. At the back or lower level away from delicate instruments, with frequently used items within easy reach.

🧠 Clinical Rationale: Layout protects delicate items from damage and places commonly used instruments where the scrub nurse can reach them instantly 🧰.

🔍 Key Clues: Heavy low, delicate safe, common close.

❌ Why Not Others: (b) stacking damages delicate instruments; (c) removal from the table defeats usability; (a) sidedness should follow the user's handedness.

⚡ 5-Second Recall: Set the table for speed and safety 🎯

📚 Ref: AORN setup guidance.

Positive Marks: 1

Negative Marks: 0.33

Question: Instruments must be kept separate from rubbish and linen throughout the case because

(a) rubbish bags are rated for instruments

(b) it is a hospital tradition

(c) accidental disposal of instruments causes retention risk and expensive loss, and instruments can injure waste handlers

(d) linen stores need space

Answer: c

Solution: ✅ Correct Answer: c. Accidental disposal of instruments causes retention risk and expensive loss, and instruments can injure waste handlers.

🧠 Clinical Rationale: Physical separation plus count discipline prevents instruments reaching waste or linen streams 🗑️.

🔍 Key Clues: Instruments stay in the instrument zone.

❌ Why Not Others: (b) tradition is not a reason; (d) linen storage is unrelated; (a) no rubbish bag is intended to carry instruments.

⚡ 5-Second Recall: Instruments never go in the bin 🎯

📚 Ref: AORN count guidance; waste handling safety.

Positive Marks: 1

Negative Marks: 0.33

Question: Before the first case of the day, the scrub nurse should verify that the instrument set

(a) contains only the instruments used in the previous case

(b) matches the tray list, is complete and functional, and the requested implants or extras are available

(c) is at room temperature only

(d) has been opened by the circulator for speed

Answer: b

Solution: ✅ Correct Answer: b. Matches the tray list, is complete and functional, and the requested implants or extras are available.

🧠 Clinical Rationale: Pre-case verification prevents mid-surgery surprises and supports the initial count 🔍.

🔍 Key Clues: Check completeness before the skin.

❌ Why Not Others: (c) temperature is not completeness; (d) premature opening risks contamination before the patient is ready; (a) current needs govern the set, not yesterday's case.

⚡ 5-Second Recall: Complete set before incision 🎯

📚 Ref: AORN setup standards.

Positive Marks: 1

Negative Marks: 0.33

Question: Metzenbaum scissors are designed for

(a) delicate dissection of fine tissue and cutting sutures, being longer and lighter than Mayo scissors

(b) cutting tough fascia and heavy tissue

(c) grasping bowel wall

(d) holding needles

Answer: a

Solution: ✅ Correct Answer: a. Delicate dissection of fine tissue and cutting sutures, being longer and lighter than Mayo scissors.

🧠 Clinical Rationale: Metzenbaum blades are long and curved for fine dissection in vascular and delicate tissue planes ✂️.

🔍 Key Clues: Metzenbaum: delicate dissection.

❌ Why Not Others: (b) Mayo scissors handle heavy tissue; (c) bowel grasping uses Babcock or Allis; (d) needles are held with needle holders.

⚡ 5-Second Recall: Metzenbaum for delicate work 🎯

📚 Ref: instrument atlases.

Positive Marks: 1

Negative Marks: 0.33

Question: Suction tips should be inspected to ensure

(a) the lumen is clear and the tip is intact, since blockage or a cracked tip fails at the critical moment

(b) the handle is brightly coloured

(c) the tubing is the longest available

(d) the tip is disposable in all cases

Answer: a

Solution: ✅ Correct Answer: a. The lumen is clear and the tip is intact, since blockage or a cracked tip fails at the critical moment.

🧠 Clinical Rationale: Suction failure during bleeding is dangerous, so patency and integrity are checked before and during use 🩸.

🔍 Key Clues: Clear lumen, intact tip.

❌ Why Not Others: (b) colour is irrelevant; (c) longer tubing is not automatically better; (d) some suction tips are reusable and need checks.

⚡ 5-Second Recall: Suction: test patency 🎯

📚 Ref: AORN equipment checks.

Positive Marks: 1

Negative Marks: 0.33

Question: A surgeon's specific instrument preference should be managed by

(a) adding extra instruments to every tray

(b) unofficial notes left in the theatre drawer

(c) a formally updated preference card agreed with the surgeon and theatre manager, with the set changes communicated to CSSD

(d) the scrub nurse's memory alone

Answer: c

Solution: ✅ Correct Answer: c. A formally updated preference card agreed with the surgeon and theatre manager, with the set changes communicated to CSSD.

🧠 Clinical Rationale: Documented preferences keep preparations accurate and traceable without ad hoc tray inflation 🗂️.

🔍 Key Clues: Preferences, written and approved.

❌ Why Not Others: (b) informal notes get lost; (d) memory fails under pressure; (a) adding to every tray wastes resources.

⚡ 5-Second Recall: Preference cards keep order 🎯

📚 Ref: theatre management guidance.

Positive Marks: 1

Negative Marks: 0.33

Question: When passing a scalpel to the surgeon, the scrub nurse should

(a) place it in a neutral zone or on the tray, or pass handle-first with the blade away from the surgeon's hand

(b) leave it loose on the drape near the surgeon

(c) throw it onto the field

(d) hold the blade toward the surgeon's palm

Answer: a

Solution: ✅ Correct Answer: a. Place it in a neutral zone or on the tray, or pass handle-first with the blade away from the surgeon's hand.

🧠 Clinical Rationale: Blade-first and hand-to-hand passing cause cuts, so neutral-zone or handle-first technique is used 🔪.

🔍 Key Clues: Handle first, blade away.

❌ Why Not Others: (d) blade-to-palm is the classic injury mechanism; (c) thrown instruments are unpredictable; (b) loose instruments on the drape injure and get lost.

⚡ 5-Second Recall: Handle first, blade away 🎯

📚 Ref: AORN sharps safety.

Positive Marks: 1

Negative Marks: 0.33

Question: A caesarean section (LSCS) set is planned to include

(a) a tonsillectomy set

(b) large abdominal instruments plus uterine haemostats, bladder retractor and cord clamps with neonatal support equipment ready outside the field

(c) only obstetric forceps and a suction tip

(d) instruments for bowel resection

Answer: b

Solution: ✅ Correct Answer: b. Large abdominal instruments plus uterine haemostats, bladder retractor and cord clamps with neonatal support equipment ready outside the field.

🧠 Clinical Rationale: LSCS needs both abdominal exposure and uterine and neonatal readiness, organised so the baby is received safely 👶.

🔍 Key Clues: LSCS: abdomen plus uterus plus baby.

❌ Why Not Others: (c) obstetric forceps are for vaginal delivery; (d) bowel instruments are not required for a routine caesarean; (a) tonsillectomy sets belong to ENT surgery.

⚡ 5-Second Recall: LSCS: mother plus newborn readiness 🎯

📚 Ref: obstetric theatre standards; instrument atlases.

Positive Marks: 1

Negative Marks: 0.33

Question: A rubber band or tape should not be used inside an instrument tray to bundle items because

(a) it harbours soil, blocks steam and breaks down in the steriliser, contaminating the set

(b) it makes the tray heavy

(c) autoclaves cannot process trays with tapes at all

(d) it changes the colour of the tray

Answer: a

Solution: ✅ Correct Answer: a. It harbours soil, blocks steam and breaks down in the steriliser, contaminating the set.

🧠 Clinical Rationale: Trays use approved holders and pins, never rubber bands, since these degrade and hold contamination 🚫.

🔍 Key Clues: No rubber bands in sets.

❌ Why Not Others: (b) weight is not the concern; (d) colour change is irrelevant; (c) the issue is contamination, not machine failure.

⚡ 5-Second Recall: Approved holders, not rubber bands 🎯

📚 Ref: CSSD preparation standards.

Positive Marks: 1

Negative Marks: 0.33

Question: Instruments with detachable parts should always be

(a) sterilised as separate unlabelled pieces

(b) packed loosely so parts do not get lost

(c) stored disassembled for convenience

(d) reassembled correctly and checked for completeness and function before packing, with all components secured

Answer: d

Solution: ✅ Correct Answer: d. Reassembled correctly and checked for completeness and function before packing, with all components secured.

🧠 Clinical Rationale: Complete, correctly assembled instruments function when needed and prevent small parts reaching the patient 🔩.

🔍 Key Clues: Complete and assembled before packing.

❌ Why Not Others: (b) loose parts get lost; (a) unlabelled pieces lose traceability; (c) disassembled storage defeats readiness.

⚡ 5-Second Recall: Whole instrument, whole set 🎯

📚 Ref: CSSD assembly standards.

Positive Marks: 1

Negative Marks: 0.33

Question: Toothed (dissecting) forceps differ from non-toothed (smooth) forceps because they

(a) are always shorter than smooth forceps

(b) grip skin and fascia securely, whereas smooth forceps handle delicate tissue without trauma

(c) are used only for bone

(d) cut tissue as they grip

Answer: b

Solution: ✅ Correct Answer: b. Grip skin and fascia securely, whereas smooth forceps handle delicate tissue without trauma.

🧠 Clinical Rationale: Teeth prevent slippage on tough tissue, and smooth jaws protect vessels, nerves and bowel 🩺.

🔍 Key Clues: Teeth for tough, smooth for delicate.

❌ Why Not Others: (c) bone uses bone-holding forceps; (d) forceps grip, they do not cut; (a) length is not the distinguishing feature.

⚡ 5-Second Recall: Toothed for skin, smooth for vessels 🎯

📚 Ref: instrument atlases.

Positive Marks: 1

Negative Marks: 0.33

Question: Instrument parts commonly named include

(a) jaws, shanks, box lock, ratchet and finger rings

(b) thread, swage and point

(c) barrel, plunger and tip

(d) blade, hub and bevel

Answer: a

Solution: ✅ Correct Answer: a. Jaws, shanks, box lock, ratchet and finger rings.

🧠 Clinical Rationale: Knowing the parts guides inspection, cleaning of locks and hinges, and correct handling 🔧.

🔍 Key Clues: Jaws, lock, ratchet, rings.

❌ Why Not Others: (d) describes a scalpel blade; (c) describes a syringe; (b) describes a suture needle.

⚡ 5-Second Recall: Know the parts to clean the parts 🎯

📚 Ref: instrument care guidance.

Positive Marks: 1

Negative Marks: 0.33

Question: When the surgeon asks for an instrument not on the tray, the scrub nurse should

(a) improvise with a different instrument silently

(b) refuse and continue with the procedure

(c) leave the field to fetch it personally

(d) ask the circulator to obtain it without leaving the field, and confirm availability before the need becomes urgent

Answer: d

Solution: ✅ Correct Answer: d. Ask the circulator to obtain it without leaving the field, and confirm availability before the need becomes urgent.

🧠 Clinical Rationale: The scrub nurse never leaves the field; the circulator retrieves items, and early confirmation avoids delays 🔄.

🔍 Key Clues: Scrub stays, circulator fetches.

❌ Why Not Others: (c) leaving the field breaks sterility and coverage; (a) silent substitution risks harm; (b) refusal without communication hampers the surgeon.

⚡ 5-Second Recall: Scrubbed hands stay put 🎯

📚 Ref: AORN role guidance.

Positive Marks: 1

Negative Marks: 0.33

Question: Cannulated instruments and screws must be cleaned with special attention to their

(a) colour coding

(b) central lumen, which must be flushed and brushed so soil and steam-blocking air are removed

(c) weight distribution

(d) packaging material

Answer: b

Solution: ✅ Correct Answer: b. Central lumen, which must be flushed and brushed so soil and steam-blocking air are removed.

🧠 Clinical Rationale: A blocked lumen both hides soil and prevents steam, so canals are flushed, brushed and dried 🔧.

🔍 Key Clues: Flush the cannulation.

❌ Why Not Others: (a) colour coding does not clean; (c) weight is irrelevant; (d) packaging comes after cleaning.

⚡ 5-Second Recall: Lumen clean, lumen sterile 🎯

📚 Ref: CSSD instrument cleaning; orthopaedic guidance.

Positive Marks: 1

Negative Marks: 0.33

Question: Instruments used on a patient with suspected prion disease require

(a) no special processing

(b) special handling with quarantine, incineration or destruction following national guidance, since routine sterilisation is insufficient

(c) the usual autoclave cycle only

(d) ordinary washing and reuse

Answer: b

Solution: ✅ Correct Answer: b. Special handling with quarantine, incineration or destruction following national guidance, since routine sterilisation is insufficient.

🧠 Clinical Rationale: Prions resist standard sterilisation, so national policies direct quarantine and destruction of instruments 🧠.

🔍 Key Clues: Prions break the usual rules.

❌ Why Not Others: (c) standard cycles do not inactivate prions; (d) reuse risks transmission; (a) special processing is mandatory.

⚡ 5-Second Recall: Prions: quarantine, destroy 🎯

📚 Ref: national prion disease guidance; CSSD policy.

Positive Marks: 1

Negative Marks: 0.33

Question: The ratchet on a haemostat exists to

(a) cut tissue

(b) attach the instrument to a tray

(c) measure the depth of a wound

(d) lock the jaws at a chosen tightness so the instrument holds without continuous hand pressure

Answer: d

Solution: ✅ Correct Answer: d. Lock the jaws at a chosen tightness so the instrument holds without continuous hand pressure.

🧠 Clinical Rationale: Ratchets free hands and maintain consistent clamping, and they must be opened for cleaning 🔒.

🔍 Key Clues: Ratchet locks the grip.

❌ Why Not Others: (a) cutting is a scissors action; (c) depth is measured by probes or rulers; (b) instruments rest in trays without attachment.

⚡ 5-Second Recall: Ratchet holds the bite 🎯

📚 Ref: instrument atlases.

Positive Marks: 1

Negative Marks: 0.33

Question: The standard scalpel handle used with a number 10 blade in general surgery is

(a) a toothed forceps

(b) a number 7 handle for all blades

(c) a needle holder

(d) a number 3 or number 4 handle depending on blade size requirement

Answer: d

Solution: ✅ Correct Answer: d. A number 3 or number 4 handle depending on blade size requirement.

🧠 Clinical Rationale: Handle and blade numbers must match, with number 3 for small blades, and firm locking of the blade before use 🔪.

🔍 Key Clues: Handle and blade must fit.

❌ Why Not Others: (b) a number 7 handle suits specific blades; (c) needle holders suture; (a) forceps grasp tissue.

⚡ 5-Second Recall: Match blade to handle 🎯

📚 Ref: instrument atlases.

Positive Marks: 1

Negative Marks: 0.33

Question: Why should instruments never be soaked in saline for long periods?

(a) Saline hardens the ratchets

(b) Saline makes instruments slippery to hold

(c) Saline corrodes stainless steel, causing pitting and rust that damage instruments and trap organisms

(d) Saline bleaches the metal

Answer: c

Solution: ✅ Correct Answer: c. Saline corrodes stainless steel, causing pitting and rust that damage instruments and trap organisms.

🧠 Clinical Rationale: Chloride ions attack the protective oxide layer, so blood and saline are rinsed off promptly 🧂.

🔍 Key Clues: Saline eats stainless steel.

❌ Why Not Others: (d) bleaching is not a saline effect; (a) hardening does not occur; (b) slipperiness is not the issue.

⚡ 5-Second Recall: No long saline soaks 🎯

📚 Ref: instrument care guidance; CSSD standards.

Positive Marks: 1

Negative Marks: 0.33