Question: Long-term exposure of theatre staff to trace concentrations of nitrous oxide is of concern mainly because it is associated with

(a) cataract formation within days of exposure

(b) reproductive harm such as reduced fertility, spontaneous abortion and neurological effects

(c) permanent loss of hearing

(d) immediate loss of consciousness

Answer: b

Solution: ✅ Correct Answer: b. Reproductive harm such as reduced fertility, spontaneous abortion and neurological effects.

🧠 Clinical Rationale: Chronic low-level nitrous oxide exposure has been linked with reproductive problems and with inactivation of vitamin B12, affecting the nervous system 🧬.

🔍 Key Clues: Waste anaesthetic gas equals staff reproductive risk.

❌ Why Not Others: (c) hearing loss is not the recognised effect of trace exposure; (a) cataracts are associated with radiation and other causes, not nitrous oxide; (d) unconsciousness requires far higher concentrations than trace leakage.

⚡ 5-Second Recall: Waste gas harms staff fertility 🎯

📚 Ref: occupational health guidance; AORN.

Positive Marks: 1

Negative Marks: 0.33

Question: During a case the return electrode alarm of the electrosurgical unit sounds. The nurse should

(a) increase the power setting to compensate

(b) press the alarm mute and continue as normal

(c) ignore the alarm if the surgery is going well

(d) ask the surgeon to stop, switch off the generator and inspect the pad and its connections

Answer: d

Solution: ✅ Correct Answer: d. Ask the surgeon to stop, switch off the generator and inspect the pad and its connections.

🧠 Clinical Rationale: The alarm signals a poor or broken pad circuit; continuing would make the current seek other paths, risking a burn at the pad or at any contact point 🚨.

🚨 Priority Principle: Safety alarms are always investigated before surgery continues.

🔍 Key Clues: Alarm equals stop, check, then resume.

❌ Why Not Others: (c) ignoring it invites a patient burn; (a) more power increases the risk without fixing the circuit; (b) muting the alarm hides the problem.

⚡ 5-Second Recall: Diathermy alarm, stop and check 🎯

📚 Ref: AORN electrosurgery safety.

Positive Marks: 1

Negative Marks: 0.33

Question: When changing an oxygen cylinder in theatre, the most important safety practice is to

(a) open the valve fully and rapidly to clear dust

(b) test for leaks with a naked flame

(c) keep oil and grease away from the fittings and open the valve slowly with the outlet facing away

(d) lubricate the valve thread to make the connection easier

Answer: c

Solution: ✅ Correct Answer: c. Keep oil and grease away from the fittings and open the valve slowly with the outlet facing away.

🧠 Clinical Rationale: Oxygen under pressure reacts violently with grease and oil and can ignite, and a rapidly opened valve produces heat and a forceful jet that can injure the operator 💨.

🚨 Priority Principle: No oil, no grease, no flame anywhere near an oxygen source.

🔍 Key Clues: Oxygen plus grease equals fire.

❌ Why Not Others: (d) lubricant is a fuel and must never touch oxygen fittings; (b) a naked flame near a leaking cylinder is extremely dangerous; (a) rapid full opening is unsafe and unnecessary.

⚡ 5-Second Recall: Oxygen: no grease, gentle opening 🎯

📚 Ref: NABH medical gas safety standards.

Positive Marks: 1

Negative Marks: 0.33

Question: During routine laparoscopy in a healthy adult, the intra-abdominal pressure is usually maintained at

(a) 40 mmHg

(b) 12 to 15 mmHg

(c) 5 mmHg

(d) 25 to 30 mmHg

Answer: b

Solution: ✅ Correct Answer: b. 12 to 15 mmHg.

🧠 Clinical Rationale: This pressure lifts the abdominal wall enough to create working space while limiting compression of the diaphragm, venous return and renal perfusion 🫁.

🔍 Key Clues: Laparoscopy plateau equals 12 to 15 mmHg.

❌ Why Not Others: (c) 5 mmHg gives inadequate exposure; (d) 25 to 30 mmHg markedly impairs venous return and ventilation; (a) 40 mmHg is dangerous and may cause gas tracking and haemodynamic collapse.

⚡ 5-Second Recall: Twelve to fifteen keeps the view without harming perfusion 🎯

📚 Ref: minimal access surgery nursing texts.

Positive Marks: 1

Negative Marks: 0.33

Question: When the defibrillator is discharged during an intra-operative arrest, the nurse must ensure that

(a) all staff stand as close as possible to the table

(b) the patient's chest is wet with alcohol

(c) nobody is touching the patient or the table and that oxygen delivery is diverted away from the paddles

(d) the diathermy remains switched on

Answer: c

Solution: ✅ Correct Answer: c. Nobody is touching the patient or the table and that oxygen delivery is diverted away from the paddles.

🧠 Clinical Rationale: A shock released into a person in contact causes injury to staff, and an oxygen-enriched field around the paddles can ignite if a spark occurs ⚡.

🚨 Priority Principle: Clear the patient and clear the oxygen before every shock.

🔍 Key Clues: Shout clear, check hands off, move oxygen.

❌ Why Not Others: (d) diathermy must be switched off to protect everyone; (b) alcohol is flammable and increases the fire risk; (a) standing near the table risks receiving the shock.

⚡ 5-Second Recall: Clear the patient, clear the oxygen, then shock 🎯

📚 Ref: advanced life support guidelines; AORN.

Positive Marks: 1

Negative Marks: 0.33

Question: The return (dispersive) electrode of the electrosurgical unit should ideally be placed

(a) directly over a metallic implant or a cardiac pacemaker

(b) over a large, well-vascularised muscle mass close to the operation site, on clean dry intact skin

(c) on a bony prominence such as the anterior shin

(d) over a scar, hairy area or previously irradiated skin

Answer: b

Solution: ✅ Correct Answer: b. Over a large, well-vascularised muscle mass close to the operation site, on clean dry intact skin.

🧠 Clinical Rationale: Current spreads out over a wide surface, so a large muscle belly with good perfusion keeps current density low and prevents heat build-up and a pad-site burn 🔥.

🔍 Key Clues: Big muscle, close to the site, clean and dry.

❌ Why Not Others: (c) bone has poor perfusion, so current concentrates and burns; (d) scarred, hairy or irradiated skin has high resistance and poor healing; (a) implants and pacemakers can channel current and cause injury.

⚡ 5-Second Recall: Muscle belly close by, never bone or scar 🎯

📚 Ref: AORN electrosurgery guidance; Miller's Anaesthesia.

Positive Marks: 1

Negative Marks: 0.33

Question: The most effective method of maintaining intra-operative normothermia in a patient having long abdominal surgery is

(a) keeping the theatre air conditioning at maximum

(b) extra drapes over the legs only

(c) giving a single blanket at the end of the procedure

(d) forced-air warming over the upper body combined with warmed intravenous and irrigation fluids

Answer: d

Solution: ✅ Correct Answer: d. Forced-air warming over the upper body combined with warmed intravenous and irrigation fluids.

🧠 Clinical Rationale: Anaesthesia causes core-to-periphery redistribution and prevents vasoconstriction, so active warming and warmed fluids are needed together to defend the core temperature 🌡.

🚨 Priority Principle: Prevent hypothermia from the start, do not treat it after extubation.

🔍 Key Clues: Active warming plus warmed fluids.

❌ Why Not Others: (b) passive covering alone cannot keep pace with heat loss; (a) cold air conditioning actively increases heat loss; (c) a blanket after the event does not prevent the earlier drop.

⚡ 5-Second Recall: Warm the patient, warm the fluids 🎯

📚 Ref: ERAS Society guidance; AORN.

Positive Marks: 1

Negative Marks: 0.33

Question: Modern operating theatre electrical circuits reduce the risk of electric shock to the patient mainly by using

(a) an isolated circuit with a line isolation monitor that alarms when leakage current rises

(b) wooden operating tables in place of metal ones

(c) high-voltage mains supply

(d) rubber-soled shoes for all staff

Answer: a

Solution: ✅ Correct Answer: a. An isolated circuit with a line isolation monitor that alarms when leakage current rises.

🧠 Clinical Rationale: Isolation means a single fault to earth does not complete a circuit through the patient, and the monitor warns staff before a second fault creates danger ⚡.

🔍 Key Clues: Isolated circuit plus monitor equals microshock protection.

❌ Why Not Others: (c) voltage is not raised for safety; (d) staff footwear protects staff, not the patient; (b) table material does not isolate the mains supply.

⚡ 5-Second Recall: Isolation and a monitor protect the patient 🎯

📚 Ref: AORN electrical safety; OT nursing texts.

Positive Marks: 1

Negative Marks: 0.33

Question: After a needle-stick injury in theatre, the first action is to

(a) wash the wound gently with soap and running water without squeezing or scrubbing the puncture

(b) apply a tight tourniquet above the injury

(c) squeeze the wound hard to make it bleed

(d) scrub the site with an alcohol swab for five minutes

Answer: a

Solution: ✅ Correct Answer: a. Wash the wound gently with soap and running water without squeezing or scrubbing the puncture.

🧠 Clinical Rationale: Gentle washing removes surface contamination; squeezing or scrubbing damages tissue and may push contaminants deeper, and caustic agents injure the wound 🧼.

🚨 Priority Principle: Wash, report, then start post-exposure prophylaxis as soon as possible.

🔍 Key Clues: Soap and water, do not squeeze.

❌ Why Not Others: (c) milking the wound does not reduce risk and causes trauma; (d) alcohol is an irritant and does not inactivate blood-borne viruses reliably; (b) a tourniquet increases venous pressure and bleeding without helping.

⚡ 5-Second Recall: Wash gently, report at once 🎯

📚 Ref: CDC post-exposure guidance; BMW Rules 2016.

Positive Marks: 1

Negative Marks: 0.33

Question: For an average adult, a pneumatic tourniquet on the arm is usually inflated to about

(a) 400 mmHg

(b) 100 mmHg, below systolic pressure

(c) 500 mmHg

(d) 250 mmHg, roughly 50 to 100 mmHg above systolic pressure

Answer: d

Solution: ✅ Correct Answer: d. 250 mmHg, roughly 50 to 100 mmHg above systolic pressure.

🧠 Clinical Rationale: The cuff must exceed systolic pressure enough to occlude arterial flow, but not so much that it crushes muscle and nerve beneath it 💪.

🔍 Key Clues: Arm cuff equals about 250 mmHg.

❌ Why Not Others: (b) a pressure below systolic leaves arterial flow intact and does not produce a bloodless field; (a) 400 mmHg causes nerve and muscle injury; (c) 500 mmHg causes severe tissue damage.

⚡ 5-Second Recall: Arm 250, leg 300 🎯

📚 Ref: tourniquet practice guidance; AORN.

Positive Marks: 1

Negative Marks: 0.33

Question: Body jewellery and metal rings should be removed before monopolar diathermy mainly because

(a) jewellery makes positioning the patient difficult

(b) they may be lost in the theatre and cause a medico-legal dispute

(c) metal items interfere with the theatre lighting

(d) current can concentrate at metal contact points and cause a burn

Answer: d

Solution: ✅ Correct Answer: d. Current can concentrate at metal contact points and cause a burn.

🧠 Clinical Rationale: Metal is an excellent conductor; a ring, chain or piercing in contact with damp skin or an earth path focuses current and generates local heat and a burn 💍.

🔍 Key Clues: Metal, damp skin, heat, burn.

❌ Why Not Others: (a) positioning is a convenience issue, not the safety reason; (c) metal does not affect theatre lighting; (b) losing valuables is a separate administrative concern.

⚡ 5-Second Recall: Any metal on damp skin can burn 🎯

📚 Ref: AORN electrosurgery guidance.

Positive Marks: 1

Negative Marks: 0.33

Question: When a patient is receiving supplemental oxygen by mask or nasal prongs, diathermy in the head and neck region is hazardous mainly because

(a) an oxygen-enriched atmosphere around the face and drapes makes ignition much more likely

(b) oxygen corrodes the diathermy pencil

(c) diathermy reduces the inspired oxygen concentration

(d) oxygen increases the electrical resistance of the skin

Answer: a

Solution: ✅ Correct Answer: a. An oxygen-enriched atmosphere around the face and drapes makes ignition much more likely.

🧠 Clinical Rationale: Even a modest rise in the oxygen fraction greatly increases the speed and intensity of combustion, so oxygen is stopped or diverted for several minutes before diathermy in an open field 🔥.

🚨 Priority Principle: Where oxygen is enriched, the ignition source must be controlled.

🔍 Key Clues: Open oxygen plus spark equals flash fire.

❌ Why Not Others: (b) oxygen does not corrode the pencil; (d) skin resistance is affected by moisture, not oxygen concentration; (c) diathermy does not alter the inspired oxygen.

⚡ 5-Second Recall: Stop the oxygen before you spark near the face 🎯

📚 Ref: AORN fire safety in head and neck surgery.

Positive Marks: 1

Negative Marks: 0.33

Question: An electrical fire involving a piece of operating theatre equipment is best extinguished with

(a) a bucket of water

(b) a dry cloth wrapped around the equipment

(c) a carbon dioxide extinguisher

(d) an alcohol-based foam

Answer: c

Solution: ✅ Correct Answer: c. A carbon dioxide extinguisher.

🧠 Clinical Rationale: Carbon dioxide is non-conducting and leaves no residue, so it extinguishes a live electrical fire without exposing staff to shock 🔌.

🔍 Key Clues: Electrical fire equals carbon dioxide.

❌ Why Not Others: (a) water conducts electricity and electrocutes the user; (d) an alcohol-based agent adds fuel; (b) a cloth smothers nothing and may catch fire and burn the hands.

⚡ 5-Second Recall: Electrical fire, carbon dioxide 🎯

📚 Ref: fire safety standards; NABH.

Positive Marks: 1

Negative Marks: 0.33

Question: A patient develops a burn over the return electrode site while diathermy is in use. The most likely cause is

(a) a deliberately low power setting on the generator

(b) a pad that has partly peeled away, so the remaining contact area carries too much current

(c) irrigation of the wound with cold saline

(d) the use of bipolar forceps during the case

Answer: b

Solution: ✅ Correct Answer: b. A pad that has partly peeled away, so the remaining contact area carries too much current.

🧠 Clinical Rationale: When only part of the pad touches skin, the same current passes through a smaller area, current density rises sharply and the skin heats up and burns ⚡.

🚨 Priority Principle: A pad-site burn always means lost contact area or poor skin contact.

🔍 Key Clues: Reduced pad area equals increased current density.

❌ Why Not Others: (d) bipolar current flows only between the forceps tips and needs no return pad; (a) a lower power setting reduces, not increases, heating; (c) cold irrigation does not heat tissue.

⚡ 5-Second Recall: Small contact area, big burn 🎯

📚 Ref: AORN electrosurgery safety; MHAUS.

Positive Marks: 1

Negative Marks: 0.33

Question: A Poole suction tip is chosen in preference to a Yankauer mainly because it

(a) is designed for oropharyngeal secretions

(b) has multiple holes that clear large volumes of blood and irrigation fluid from a body cavity without blocking

(c) is used to instil fluid into the abdominal cavity

(d) is the correct tip for tracheal suctioning

Answer: b

Solution: ✅ Correct Answer: b. Has multiple holes that clear large volumes of blood and irrigation fluid from a body cavity without blocking.

🧠 Clinical Rationale: The pool tip draws fluid from many apertures at once, so it moves a large volume quickly and resists clogging by clot or tissue 💧.

🔍 Key Clues: Poole equals pool of fluid in a cavity.

❌ Why Not Others: (a) oropharyngeal clearance is the role of the Yankauer; (c) suction tips remove rather than instil fluid; (d) tracheal suction uses a sterile single-lumen catheter.

⚡ 5-Second Recall: Poole for pools, Yankauer for the throat 🎯

📚 Ref: AORN equipment standards; OT nursing texts.

Positive Marks: 1

Negative Marks: 0.33

Question: Surgical smoke generated by a laser or diathermy is best controlled by

(a) opening the theatre doors during the procedure

(b) an ordinary wall suction tip placed a metre away from the field

(c) a dedicated smoke evacuator with the nozzle held close to the site of plume generation

(d) allowing the smoke to disperse naturally in the theatre

Answer: c

Solution: ✅ Correct Answer: c. A dedicated smoke evacuator with the nozzle held close to the site of plume generation.

🧠 Clinical Rationale: Smoke plume carries viable micro-organisms, toxic chemicals and carcinogens, and capture falls off rapidly with distance so the nozzle must be at the source 🌫.

🚨 Priority Principle: No staff member should breathe visible plume.

🔍 Key Clues: Capture at the source, not at the ceiling.

❌ Why Not Others: (b) ordinary suction filters poorly and loses capture with distance; (a) open doors disrupt air changes and do not remove plume from the breathing zone; (d) passive dispersal exposes everyone in the room.

⚡ 5-Second Recall: Suck the smoke at the source 🎯

📚 Ref: AORN smoke evacuation guidance.

Positive Marks: 1

Negative Marks: 0.33

Question: When the electrosurgical pencil is not being used, it should be kept

(a) in an insulated holster, with the tip pointed away from the patient

(b) on the sterile drapes beside the wound

(c) hanging over the side of the operating table

(d) resting on the suction tubing next to the skin

Answer: a

Solution: ✅ Correct Answer: a. In an insulated holster, with the tip pointed away from the patient.

🧠 Clinical Rationale: An activated pencil left on drapes or skin can burn the patient or start a drape fire, so a holster keeps the hot tip isolated 🧯.

🔍 Key Clues: Idle pencil lives in a holster.

❌ Why Not Others: (b) drapes are flammable and can ignite; (c) it may fall and injure someone or activate against the patient; (d) tubing offers no insulation and lies in contact with the patient.

⚡ 5-Second Recall: Pencil down means pencil in the holster 🎯

📚 Ref: AORN fire and electrosurgery safety.

Positive Marks: 1

Negative Marks: 0.33

Question: During insufflation a patient suddenly becomes hypotensive with a sharp fall in end-tidal carbon dioxide and a mill wheel murmur. The most likely cause is

(a) a carbon dioxide gas embolism

(b) a full bladder

(c) light anaesthesia

(d) an allergic reaction to the suture material

Answer: a

Solution: ✅ Correct Answer: a. A carbon dioxide gas embolism.

🧠 Clinical Rationale: Gas in the pulmonary circulation obstructs outflow, so end-tidal carbon dioxide collapses and blood pressure falls, while the classic mill wheel murmur may be heard 🚨.

🚨 Priority Principle: Sudden hypotension with falling end-tidal carbon dioxide means stop insufflation and tell the surgeon and anaesthetist at once.

🔍 Key Clues: Mill wheel murmur, falling end-tidal carbon dioxide.

❌ Why Not Others: (c) light anaesthesia raises rather than collapses end-tidal carbon dioxide and pressure; (b) a full bladder does not produce a murmur; (d) a suture allergy does not cause an abrupt haemodynamic collapse during insufflation.

⚡ 5-Second Recall: Gas embolus means stop the gas and call out 🎯

📚 Ref: Miller's Anaesthesia; AORN.

Positive Marks: 1

Negative Marks: 0.33

Question: Instruments used in a laser field should have a matt, non-reflective finish because

(a) non-reflective instruments conduct electricity better

(b) reflective instruments are heavier and tire the surgeon

(c) a shiny surface can reflect the beam unpredictably and cause an injury away from the target

(d) matt instruments are cheaper to sterilise

Answer: c

Solution: ✅ Correct Answer: c. A shiny surface can reflect the beam unpredictably and cause an injury away from the target.

🧠 Clinical Rationale: A reflected beam retains enough energy to burn tissue, drapes or eyes elsewhere in the room, so matt black instruments and retractors are used in the field 🪞.

🔍 Key Clues: No mirrors in a laser field.

❌ Why Not Others: (d) cost and sterilisation are unrelated to beam safety; (b) the weight of the instrument is not the reason; (a) laser safety concerns light, not electrical conduction.

⚡ 5-Second Recall: Matt instruments, no reflections 🎯

📚 Ref: AORN laser safety.

Positive Marks: 1

Negative Marks: 0.33

Question: In a patient with a previous midline laparotomy, the safest initial entry point for insufflation is usually

(a) the right lower quadrant over the appendix scar

(b) the umbilicus, as in every other patient

(c) through the largest existing scar

(d) an alternative site in the left upper quadrant, such as Palmer's point

Answer: d

Solution: ✅ Correct Answer: d. An alternative site in the left upper quadrant, such as Palmer's point.

🧠 Clinical Rationale: Adhesions form most densely beneath an old midline or lower abdominal scar, so an entry point away from scarred bowel reduces the risk of visceral injury 🩺.

🔍 Key Clues: Previous surgery equals consider an alternative entry point.

❌ Why Not Others: (b) the umbilicus is directly below the midline scar and carries the highest adhesion risk; (a) a quadrant over an old scar is just as unsafe; (c) a large scar is the very place where adherent bowel is likely.

⚡ 5-Second Recall: Old midline scar means choose another point 🎯

📚 Ref: minimal access surgery texts.

Positive Marks: 1

Negative Marks: 0.33